## LEISTUNGSVERZEICHNIS
LEISTUNGSVERZEICHNIS / Tender Specification |  |  |  |  | Stand: 31.03.2026
Banf-Nummer/order request no.: |  | PR1270418
Anzahl/quantity: |  | 1 Piece
Gerätesteckbriefnummer/device profile no: |  | IMWS-04.2 & IMWS-04.3
Auftragsgegenstand/order subject: |  | High resolution X-ray inspection tool
KURZE BESCHREIBUNG DER LEISTUNG/ BRIEF SERVICE DESCRIPTION:
The Fraunhofer Institute for Microstructure of Materials and Systems IMWS researches the application behaviour, reliability, safety and lifetime of innovative materials in components and systems. The Business Unit “Electronic Materials and Components” Fraunhofer IMWS support industry and research partners by its internationally recognized expertise in failure analyses of electronic components. Fraunhofer IMWS will enhance its capabilities for advanced failure analysis and material characterization to support the ramp up of new heterogenous integration technologies addressing the manufacturing and reliability challenges. To enable the non-destructive detection and evaluation of structural deviations, contact defects, damage characteristics and their most precise possible localization in the package to enable further high-resolution physical diagnostics, the latest generation of radiography (2D) and X-ray tomography (3D-CT) systems are required in particular for the non-destructive detection and evaluation of structural deviations, contact defects, damage characteristics and their most precise possible localization in the package to enable further high-resolution physical diagnostics. A high-performance multi-processor workstation is required for postprocessing of the generated imaging data sets to identify tiny deviations and defects within the devices. This tender is for the supply a functional complete device including assembly, commissioning and all necessary connections, approvals/certifications and software for operation. The requirements for a state-of-the-art x-ray inspection tool for highly integrated 2.5/3D components are described below:
Legende Kriterien/legend criteria
B | Bewertung: Leistungsverzeichnispositionen, die mit einem B gekennzeichnet sind, dienen der Bewertung einzelner Kriterien. Die maximal erreichbare Punktzahl für das einzelne Kriterium ist in Spalte G aufgelistet. Eine Nichterfüllung oder keine Angaben führen zum sofortigen Ausschluss. Evaluation: The positions of service specification marked with B are used to evaluate individual criteria. The maximum achievable score for each criterion is listed in column J. Non-compliance or no information will result in immediate exclusion.
F | Forderung: Leistungsverzeichnispositionen, die mit einem F gekennzeichnet sind, müssen zwingend erfüllt werden. Das bedeutet für den Bieter, dass im Angebot dazu Angaben zu erstellen sind. Eine Nichterfüllung oder keine Angaben führen zum sofortigen Ausschluss. Requirement: The positions of service specification marked with F must be fulfilled. The information must be provided in the offer. Non-compliance or no information will result in immediate exclusion.
O | Option: Leistungsverzeichnispositionen, die mit einem O gekennzeichnet sind, dienen ebenfalls der Bewertung einzelner Kriterien. Die maximal erreichbare Punktzahl für das einzelne Kriterium ist in Spalte G aufgelistet. Optionen nicht anzubieten, führt jedoch nicht zum Ausschluss. Eine nicht angebotene Option wird mit 0 Punkten bewertet. Option: Specification items marked with an “O” are also used to evaluate individual criteria. The maximum number of points available for each criterion is listed in column G. However, failure to offer an option does not result in disqualification. An option that is not offered is awarded 0 points.
I | Information: Leistungsverzeichnispositionen, die mit einem I gekennzeichnet sind, fließen nicht in die Wertung ein. Diese dienen der Übermittlung von Informationen. Diese Positionen müssen zwingend beachtet werden. Sie führen nicht zum Ausschluss. Information: The positions of service specification marked with I are not included in the evaluation. They are used for the transmission of information, which must be observed and do not result in exclusion.
N | Nachhaltigkeit: Leistungsverzeichnispositionen, die mit einem N gekennzeichnet sind, dienen der Bewertung der Nachhaltigkeit als eigenes Zuschlagskriterium. Eine Nichterfüllung oder keine Angaben führen zum sofortigen Ausschluss. Sustainability: The positions of service specification marked with an N are used to assess sustainability as an award criteria. Non-compliance or no information will result in immediate exclusion.
LV Pos. / TS Pos. | TECHNISCHE BESCHREIBUNG - Hauptpunkt Technical description - main topic | TECHNISCHE BESCHREIBUNG - Beschreibung Technical description - specification | Was wird gewünscht - Parameter What is needed - parameter | B/F/O/I | max. Punktzahl B/O-Kriterien max. total points B / O-criteria | vom Bieter auszufüllen / to be filled by tenderer | Verweis in Bieterdokument / Reference in tenderer document
1. | General Device Specifications IMWS-04.2
1.1.1. | General Device Specifications | X-ray radiation protection requirement | X-ray CT system as a stand-alone device must meet the technical requirements and criteria for type approval of full-body protection devices pursuant to §21 of the German Radiation Protection Ordinance (StrlSchV). | F
1.1.2. | General Device Specifications | Space requirements | Maximum available floor space for the system, including access areas for operator, service and maintenance: 4000 mm * 3200 mm | F
1.2.1. | Technical Specifications | xray source | X-ray source suitable for high-resolution micro-CT imaging with a serviceable design allowing replacement of wear components (e.g. cathode filament) required for normal operation. | F
1.2.2. | Technical Specifications | xray source | max. accelerating voltage: ≥ 160 kV | F
1.2.3. | Technical Specifications | Power by using 160 kV (or higher) source | Output power of at least 10 W at an accelerating voltage of 160 kV or higher. | F
1.2.4. | Technical Specifications | detector-size | min. 150mm * 150mm | F
1.2.5. | Technical Specifications | detector pixel size | ≤ 150 µm | F
1.2.6. | Technical Specifications | source-detector-orientation | vertical | F
1.2.7. | Technical Specifications | total detector tilt (in two directions from source-detector-axis) | min. 60° in two directions from source-detector-axis | F
1.2.8. | Technical Specifications | Focus-Detector -Distance | variable distance between xray-source and detector (adjustable via software) | F
1.2.9. | Technical Specifications | Sample stage | Sample stage rotatable 360° in the x-y plane | F
1.3.1. | Samples | Weight | possible sample weight: ≥ 2 kg | F
1.3.2. | Samples | Size of sample support plate | min. 400 * 400 mm² | F
1.3.3. | Samples | sample safety | Dose monitoring system to avoid sample degradation by radiation | F
1.3.3. | Samples | sample safety | Possibility of dose reduction using a collimator aperture and defined material filters. | F
1.4.1. | Targets | Target materials | Target material suitable to achieve the specified spatial resolution and image quality over the required operating range. | F
1.5.1. | Resolution | maximum 2D resolution | Maximum spatial resolution ≤ 600 nm, determined using the manufacturer's documented measurement method or an internationally recognized test procedure. | F
1.5.2. | Resolution | Voxel size | max. voxel resolution of 3D-CT ≤ 1 µm | F
1.6.1. | 2.5/3D-Scanning | Micro-CT | 3D volume creation by full-rotation of the sample using rotation manipulator. Software required for acquisition, reconstruction and visualization shall be included. | F
1.6.2. | 2.5/3D-Scanning | Laminography | 3D volume creation by combination of rotating sample stage and detector tilt; Software required for acquisition, reconstruction and visualization of laminography data shall be included. | F
1.7. | General Device Specifications IMWS-04.3
1.7.1. | Reconstruction | High-Power Workstation | Main facts: - Processor with at least 36 physical CPU cores (or equivalent computing performance) suitable for high-performance 3D reconstruction. - min. 512GB RAM - min. 48GB GPU - High-performance NVMe SSD storage with at least 15 TB usable capacity. - Windows 11 | F
1.7.2. | Reconstruction | Software package | The offer must include the most comprehensive total package available from the respective software manufacturer as a single-node license, including all software modules available at the time of delivery for volume reconstruction, visualization, analysis, evaluation, and reporting. The license must include lifetime free updates and upgrades. | F
2. | Safety and Compliance
2.1. | Safety and Compliance | According to machine guidelines the system has to include safety features to protect the operator and equipment. Emergency Stop Buttons: Easily accessible emergency stop mechanisms. Interlock Switches: To prevent unauthorized or unsafe operation. Light indicator: Adequate lighting while device is under operation, conforming to safety standards. |  | F
3. | Other Specifications
3.1. | Other Specifications | Initial user training (on-site; at least two days) for operating and data analysis is included. |  | F
3.2. | Other Specifications | The supplier shall offer advanced user training (minimum three days) |  | F
3.3. | Other Specifications | Costs for maintenance and service contracts | Please list all types/levels of maintenance and service contracts and specify the respective scope of services. Please state the annual costs in euros. | I
3.4. | Other Specifications | The supplier shall provide the tools and consumables required for initial commissioning, acceptance testing and initial sample measurements. |  | F
3.5. | Other Specifications | Technical documentation,, instructions, device labels, in German or English language. |  | F
4. | Delivery and Acceptance
4.1. | Delivery and Acceptance | Transport, packaging, shipping, customs and costs for import of dual use goods of all components and accessories have to be included. Delivery and Installation address: Fraunhofer IMWS Heideallee 19 (adjoining building) 06120 Halle, ground floor (DPU) |  | I
4.2. | Delivery and Acceptance | For the acceptance of the device, a protocol is drawn up together with the supplier in which the required parameters are recorded and checked. |  | I
4.3. | Delivery and Acceptance | Please ensure that the devices are brought to the appropriate laboratory. Transport equipment etc. must be brought with you, as we cannot store it on site. Please inform yourself about all relevant requirements with the technical contact person in good time. |  | I
5. | sustainability
5.1. |  | The main manufacturing activities and the assembly of the equipment shall be performed within Europe in order to minimize transportation distances and associated environmental impacts. |  | F
| Pricing | 90%
| Substainability | 10%

## Anhänge zum LV
Anhänge zum LV /Appendices to the LV
Banf-Nummer/order request no.: |  | PR1270418
Anzahl/quantity: |  | 1 Piece
Gerätesteckbriefnummer/device profile no: |  | IMWS-04.2 & IMWS-04.3
Auftragsgegenstand/order subject: |  | High resolution X-ray inspection tool
LV Pos. / TS Pos. | TECHNISCHE BESCHREIBUNG - Hauptpunkt Technical description - main topic | Anhänge zum LV /Appendices to the LV
1.1 | Überschrift Thema 1 Headline Topic 1 | Anhang