|  |  | MEDICAL DEVICE SPECIFICATION |
| --- | --- | --- |
|  | Name | Full Field Digital Mammography System |
|  | Quantity | 1 |
|  |  | TECHNICAL CHARACTERISTICS |
| 1 | Detailed requirements | A fully digital mammography system designed for high-quality breast imaging, suitable for screening and patient follow-up in a Primary Health Center (PHC) setting, acting as satellite of a secondary hospital for this service. - The system should be configured and supplied for 2D full-field digital mammography (FFDM) use. - While offered systems may be technically capable of digital breast tomosynthesis (DBT / 3D mammography), tomosynthesis hardware, software licenses, reconstruction algorithms, and clinical applications are not required under this procurement. The system architecture may support future upgrade to DBT / 3D mammography, subject to separate procurement and regulatory approval. The system shall provide: - High-resolution 2D FFDM imaging - Low radiation dose, optimized through advanced AEC - Workflow efficiency and patient comfort - PACS/RIS/HIS integration (DICOM 3 compliant) - State-of-the-art performance consistent with leading international manufacturers 2. System Configuration - Complete Digital mammography imaging unit (gantry, tube, generator, detector, etc). - One (1) Acquisition Workstation / Operator Console - One (1) Review / Evaluation Workstation (RWS) - Image processing and evaluation software specific to mammography. X-ray Tube and Gantry Assembly, - Motorized Compression System with interchangeable paddles, - Radiation protection panel: Suitable X-ray shielding barrier for operator safety, mobile type as per room/device design. II. Uninterruptible Power Supply (UPS) - Manufacturer-recommended rating to support safe workflow completion and controlled shutdown during power interruptions. III. Associated Safety Equipment - Operator radiation shield panel, as mentioned above. - Apron 3 Other features. 3.1 X-ray Generator - Power output: ≥ 5 kW (continuous rating), suitable for high-throughput FFDM operation. - Tube voltage range: 20 – 49/50 kVp, increment steps ≤ 1 kVp - Tube current load range: ≥ 5 – 400 mAs - Exposure control: Fully automatic with manual override - Manual and automated selection of kVp, mAs, filter - Dual-track target required (e.g., W/Rh, W/Ag, or equivalent) - AEC: multi-zone, adaptive, tissue-sensing algorithm. - AEC dose reproducibility: ≤ ±5% (per IEC 61223-3-2). 3.2 X-ray Tube - Tube voltage range, approx.: 20 – 40 kVp, adjustable - Maximum power: approx. 6–8 kW - Selectable filters: Mo, Rh, Ag, Al (or equivalent combinations) - Dual focal spots with manual/automatic selection - Anode heat capacity ~300 kHU or manufacturer-equivalent thermal performance; tube unit heat storage to support continuous workflow. - Anode material: Tungsten - Capable of delivering low dose / high penetration spectra, for dense breast imaging 3.3 Gantry / Tube Head - Fully automated collimation adapted to different paddle sizes. - Motorized vertical and rotational movements - Vertical speed: approx. 50 mm/s - Rotation range: ≥ 270° isocentric - Rotational speed: approx. 10°/s - Positioning control buttons on both sides of the arm - Automatic positioning presets modes: Yes - SID ≥ 65 cm - Safety features: - Movement limitation when compression force exceeds set threshold compression force - Collision protection system 3.4 Compression System - Motion range: approx. 0 – 180 mm - Compression force, adjustable: approx. 0 – 200 N. - Compression accuracy: ±1 daN - Paddle tilt reproducibility: ±1° - Automatic and manual modes - Smooth, motorized application with emergency/manual release 3.5 Exposure Control - Modes: Automatic (AEC) and manual Accuracy: kVp indicator ± 2 % mAs indicator ± 10 % - Exposure time appropriate for 2D FFDM imaging, approx.: < 2 seconds. - Exposure parameters optimized for full-field digital mammography (FFDM) 3.6 Detector and Imaging - Direct digital flat-panel detector. - Pixel size ≤ 100 µm, with detector performance in accordance with IEC 62220-1-2. - Detector technology based on direct or indirect conversion, with high quantum efficiency suitable for mammography. - Active imaging area, approx. ≥ 24 × 29/30 cm. - DQE performance demonstrated in accordance with IEC 62220-1-2 with DQE values consistent with current clinical mammography systems. - Native spatial resolution ≥ 19 lp/mm. - Integrated anti-scatter grid suitable for mammography, high-transmission design; removable or appropriately configurable for magnification views, as applicable. - Detector uniformity and image quality compliance with IEC 61223 acceptance and constancy tests. 3.7 Acquisition Workstation (AWS). - (1) one medical-grade grayscale diagnostic monitor (≥21.3”, ≥5 MP, DICOM GSDF compliant) for image acquisition and review, and (1 ) one secondary monitor (≥19”, ≥2 MP) for system control and workflow management. - Real-time acquisition and preview of images - Full DICOM compliance (Store, Query/Retrieve, MPPS, Print) - Local image review: zoom, pan, window/level, multi-frame display - Integrated patient registration and exam protocol interface - Full control of generator and detector parameters 3.8 Workstation Hardware (minimum): - Processor: Intel Core i7 or AMD Ryzen 7, or higher - Memory: ≥ 16 GB RAM - Storage: ≥ 1 TB HDD + ≥ 512 GB SSD (approx. 10,000 images) - Monitor: Two (2 x) medical-grade diagnostic monitor ≥ 21.3 Inches, 5 MP. - Black and white, high resolution, high brightness, flicker-free and anti-glare, capable of displaying various formats, with brightness at least 500 cd/m2. - Input: Medical-grade keyboard and mouse - Connectivity: USB 3.0, Ethernet, HDMI/DP; RS232 / LAN interface / DICOM 3.0 GSDF compliant - Power backup: Integrated UPS, capacity appropriately sized to support safe operation and controlled shutdown of the workstation. - Exposure settings: Manual and fully automatic. - Suitable for primary mammographic interpretation. - Multimodality capable of accessing patient images obtained from other modalities such as MRI, CT, X-ray, or Ultrasound. - Evaluation diagnosis software specific to mammography must be provided 3.9 Image Processing Requirements: The system should include advanced image processing, such as: - Adaptive contrast enhancement - Noise reduction algorithms - Tissue equalization - Automatic breast boundary detection - Automated pectoral muscle recognition for MLO views |
| 2 | Displayed parameters | - Real-time display of kVp, mAs - Patient demographic data (DICOM compliant) - Preview image available within ~5 s; full image available within ~10 s - Compression force - Breast thickness - Exposure time - Dose indicator (AGD/ESD or equivalent) - System status, warnings, and alarms |
| 3 | User adjustable settings | All operational parameters, among them. - Compression force and speed - Pre-set imaging protocols by type (screening/diagnostic) - Brightness/contrast and image post-processing filters - Exposure parameters |
|  |  | UTILITY REQUIREMENTS |
| 4 | Electrical, water and/or gas supply | 220–240 VAC, 50 Hz. Three-phase supply (380–400 VAC) shall be acceptable if required by the manufacturer. Side cabinet refrigeration: The X-ray generator and associated electronics shall be in a side cabinet and shall use an integrated forced-air cooling system. |
|  |  | ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS |
| 5 | Accessories | Radiation protection panel: Suitable X-ray shielding barrier for operator safety, mobile type. Uninterruptible Power Supply (UPS): Capacity to support safe system operation and controlled shutdown in case of power failure. Compression paddles: At minimum, one small and one standard size; transparent, radiolucent, and compatible with the system. Magnification platform: Radiolucent, suitable for at least one magnification factor (e.g., 1.5× or 2.0×). Footswitch or remote control: For initiating exposure and/or compression release. Patient support stool: Radiolucent, height-adjustable, suitable for seated examinations. |
| 6 | Consumables/ reagents | N. A |
| 7 | Spare parts | N. A |
|  |  | TRAINING, INSTALLATION AND UTILISATION |
| 8 | Pre-installation requirements | - Pre-installation Requirements: The awarded supplier shall provide a complete and detailed set of site preparation requirements for the mammography system. This must include, at minimum: - Room layout and dimensional requirements, showing equipment footprint, workflow clearances, operator position, and patient positioning. - X-ray shielding calculations and specifications, prepared in accordance with IEC standards and Turkish radiation protection regulations, including wall/door/window requirements. - Power requirements, with clear indication of: - Electrical load (kW / Amps), - Voltage and frequency, - Dedicated line specifications, - Grounding required - The exact location and connection point for the Main Distribution Board (MDB). - Environmental conditions: required temperature, humidity, ventilation, and HVAC capacity for proper operation of equipment and side cabinet cooling. - Other utilities: confirmation that no water or medical gas supply is required. - See as reference Annex 1 room layout. - Final room layout drawings, to be delivered in both PDF and editable CAD formats, showing equipment positioning, operator shield location, and all utility points. - Responsibilities: - All civil, structural, and utility preparations (room construction, shielding installation, HVAC, electrical conduit up to the MDB) will be executed by the GIZ-contracted construction company, in accordance with the requirements provided by the supplier. - From the MDB connection point onward, all electrical installation, cabling, and integration works shall be the sole responsibility of the awarded bidder. - All documentation shall be submitted by the supplier and approved by the end-user’s technical team prior to delivery of the system. - Floor loading requirement ≤ 500 kg/m² |
| 9 | Requirements for commissioning | Authorized personnel: All installation, calibration, and commissioning activities must be performed exclusively by manufacturer/supplier certified and authorized personnel. Materials and tools: The supplier shall provide all required materials, tools, and accessories necessary for commissioning. Tests to be performed (minimum): - Image quality tests in accordance with international standards and manufacturer’s protocols. - Radiation safety verification to confirm compliance with IEC, ISO, and local regulatory requirements. - DICOM integration and connectivity testing with PACS/RIS/HIS. - User acceptance testing (UAT) carried out jointly with the end-user to verify that all functional, performance, and safety requirements are met. Documentation: Commissioning results shall be formally documented, signed by supplier and end-user representatives, and provided as part of the handover package. |
| 10 | Training of users | The supplier shall provide a structured training program for all relevant personnel. - Target: radiographers, radiologists, biomedical engineers - Duration: 1 day (8 hours) with practical hands-on sessions. - Location: on site, Malatya PHC (Primary Health Centre). - Trainer: supplier/manufacturer-authorized specialist - Language: Turkish language - Number of participants: about 4 - Minimum training modules: - System operation: Daily workflow, patient positioning, acquisition protocols, compression system handling, and safety features. - Image quality and optimization: Use of automatic exposure control, protocol selection, post-processing, and quality assurance checks. - Radiation protection: Safe use of the equipment in line with ALARA principles and local regulations. - DICOM connectivity and workflow integration: Image storage, retrieval, and transfer to PACS/RIS/HIS. - Basic troubleshooting: Routine checks, error messages, safe shutdown/start-up procedures. - Preventive maintenance awareness: Daily/weekly user care, cleaning, and basic QC with supplied phantoms. - User manuals and signed attendance/competence records must be provided. - Training completion certificates shall be issued to all participants. |
| 11 | User care | Routine cleaning and disinfection using standard hospital-grade disinfectants, with smooth, accessible surfaces and materials resistant to commonly used cleaning agents. |
|  |  | WARRANTY AND MAINTENANCE |
| 12 | Warranty | Minimum two (2) years full warranty from commissioning date according to GIZ General Purchase Conditions 2023. |
| 13 | Software / Hardware upgrade availability | Yes |
|  |  | DOCUMENTATION |
| 14 | Documentation requirements | The Supplier shall provide, at delivery and commissioning, complete technical and regulatory documentation in compliance with applicable international standards and Turkish regulatory requirements. At minimum, the following shall be supplied: 1. Manuals - User, installation, and service manuals in Turkish (English acceptable in addition). 2. Certificates and Regulatory Documentation - Valid CE certificate and Declaration of Conformity. - Evidence about the manufacturer of compliance with applicable standards (e.g., IEC 60601-1, IEC 60601-2-45, ISO 13485). - DICOM Conformance Statement. - Required approvals or registrations in accordance with Turkish Medicines and Medical Devices Authority (TMMDA) regulations. 3. Maintenance and Calibration - Manufacturer-recommended preventive maintenance schedule. - Calibration and test schedule necessary for routine service. There is not any PM (Preventive Maintenance) as part of this tender. 4. Spare Parts and Service Information - List of essential spare parts and accessories with part numbers. - Contact details of authorized service provider in Turkey. 5. Warranty letter - Two (2) year full warranty (from commissioning date according to GIZ General Purchase Conditions 2023.), in accordance with Section 12 of this Technical Specification, including defined service response times. |
|  |  | SAFETY AND STANDARDS |
| 15 | Risk Classification | Class IIb |
| 16 | Regulatory Approval | Valid CE certificate Registration with TMMDA (Turkish Medicines and Medical Devices Authority) |
| 17 | Certification International standards | Manufacturer ISO 13485:2016 – Medical devices: Quality management systems or equivalent. For the device - IEC 60601-1 (latest edition) – General requirements for basic safety and essential performance. - IEC 60601-1-2 (latest edition) – Electromagnetic compatibility (EMC). - IEC 60601-1-3 (latest edition) – Radiation protection in diagnostic X-ray equipment. - IEC 60601-2-45 (latest edition) – Particular requirements for the safety of mammographic X-ray equipment. - DICOM 3.0 compliance |
| 18 | Local representation | The bidder shall have a local representative in Turkey and can carry out technical support and aftersales service for the delivered products. Please provide the address and contact details of the local company. |

Annex 1. Room layout