## Tabelle1
Minimum required measurement tools for conducting energy audits in industry, buildings and residential buildings
№ | Name of measuring instrument
General requirements (apply to all items below): • Type/scope of procurement: Measuring instruments for field energy audits of residential buildings and apartments, public buildings (e.g. schools, hospitals, administrative buildings), and industrial enterprises / utility systems in Uzbekistan. • Purpose of use: The equipment shall enable examination of electricity consumption and power quality, combustion processes, thermal performance and heat losses of building envelopes, indoor climate / ventilation / lighting conditions, and selected industrial process parameters such as flow, pressure, vibration, rotational speed, and material thickness. • The parameters below are minimum requirements. • All instruments shall be new, unused, CE-marked (or equivalent), and professionally manufactured. • Delivery includes all necessary accessories for intended use (sensors, probes, clamps, cables, batteries / charger, carrying case, software where applicable, and manual in English). • Calibration / verification: Each instrument shall be delivered calibrated and ready for use. Calibration shall be performed by the manufacturer, an authorised service centre, or a laboratory traceable to ISO/IEC 17025 (or equivalent). The calibration certificate shall be issued within 12 months before delivery, indicate the serial number, and cover the main measured parameters / channels of the device. • Protection rating minimum IP40 for all handheld instruments unless stated otherwise below. • Submit with each item: manufacturer datasheet, CE declaration, list of included accessories, warranty confirmation, and calibration certificate. • Required evidence with offers: Manufacturer datasheet confirming compliance with all listed parameters.
1 | Vibration Analyzer * | • Frequency range: 0.5–10,000 Hz (or wider) • Acceleration ≥ 200 m/s²; velocity ≥ 200 mm/s • Channels: ≥ 2 vibration + ≥ 1 tachometer • FFT spectrum: ≥ 1,600 lines • ISO 20816 or ISO 10816 evaluation function • PC software for analysis included • Calibration certificate shall cover vibration measurement channels and tachometer input
2 | Laser tachometer* | • Range: 5–99,999 RPM (non-contact) • Accuracy: ± 0.05% of reading or better • Resolution: 0.1 RPM • Laser Class 2 per IEC 60825-1 • Reflective tape included • Delivered calibrated for rotational speed measurement
3 | Power and quality analyzer of electricity | • 3-phase + neutral; ≥ 4 voltage + ≥ 4 current inputs • Voltage up to 1,000 V; current 0.1–3,000 A (clamps/probes supplied) • IEC 61000-4-30 Class A or Class S • Harmonics to ≥ 40th order; THD • Safety: CAT III 1,000 V / CAT IV 600 V • Continuous logging ≥ 7 days; PC software included • Calibration certificate shall cover voltage inputs, current inputs, and supplied clamps / probes as a complete set
4 | Gas Analyzer** | • Gases: O₂ (0–25%), CO (0–4,000 ppm), CO₂ (0–20%), NO or NO₂ • Flue gas temperature up to +1,000 °C • Draft/pressure measurement included • Calculates: combustion efficiency, excess air (λ); selectable fuels • EN 50379 or equivalent • Built-in pump with condensate trap; PC software included • Delivered calibrated for gas channels and pressure measurement; any accessories required by the manufacturer for zero / functional check shall be included
5 | Multimeter Clamp Meter** | • True RMS; AC/DC current 0.1–1,000 A; AC/DC voltage 0–1,000 V • Accuracy: current ± 2%; voltage ± 1% • Safety: CAT III 1,000 V / CAT IV 600 V • Resistance, continuity, diode, frequency functions • Delivered calibrated for voltage and current measurement
6 | Thermal Imager** | • IR resolution: ≥ 320 × 240 pixels (≥ 640 × 480 preferred) • NETD: ≤ 50 mK • Range: −20 °C to +650 °C (or wider); accuracy ± 2 °C or ± 2% • Emissivity adjustable 0.01–1.00 • Spectral range: 7.5–14 µm • Integrated visual camera with IR overlay • IP54; drop resistance ≥ 2 m • Reporting software included • Delivered with factory calibration / verification for temperature measurement
7 | Heat flow meter | • For in-situ U-value measurement per ISO 9869-1 • Heat flux range ≥ ± 2,000 W/m²; accuracy ≤ ± 5% • Temperature sensors: ≥ 2 pcs; range −30 to +60 °C; accuracy ± 0.2 °C • Data logger: ≥ 72 h battery; selectable interval 1–60 min • U-value calculation software included • Calibration certificate shall cover heat flux sensor(s) and supplied temperature sensors
8 | Digital non-contact thermometer** | • Range: −50 °C to +800 °C (or wider) • Accuracy: ± 1.5 °C or ± 1.5% • Distance-to-spot ratio: ≥ 20:1 • Emissivity adjustable 0.10–1.00 • Laser pointer Class 2 • Delivered with factory calibration / verification for temperature measurement
9 | Laser Rangefinder** | • Range: 0.05–100 m (or wider) • Accuracy: ± 1.5 mm or better • Functions: distance, area, volume, Pythagoras • Laser Class 2; IP54
10 | Anemometer | • Velocity: 0.3–30 m/s (or wider); accuracy ± (3% + 0.1 m/s) or better • Temperature measurement: −20 to +60 °C; accuracy ± 0.5 °C • Volume flow calculation with duct dimension input • Probe suitable for duct insertion (≥ 200 mm) • Delivered calibrated for air velocity and temperature measurement
11 | Hygrometer** | • Relative humidity: 0–100% RH; accuracy ± 2% RH (10–90% range) • Temperature: −20 to +60 °C; accuracy ± 0.5 °C • Dew point display: required • Data logging with export • Delivered calibrated for humidity and temperature measurement
12 | Lux meter** | • Range: 0–200,000 lux (or wider); accuracy ± 3% or better • V(λ) corrected; cosine-corrected sensor • Detachable sensor with cable (≥ 1 m) • Delivered with calibration / verification for illuminance measurement
13 | Manometer | • Range: 0–6 MPa; accuracy ≤ 0.5% FS • Overload protection ≥ 1.5× full scale • Standard thread connection (G ¼"/G ½" or NPT) • Compatible with water, steam, compressed air, oil • Delivered calibrated for pressure measurement
14 | Ultrasonic liquid flow meter | • Clamp-on, transit-time principle • Pipe diameter: 25–1,200 mm (transducer sets supplied) • Accuracy: ± 2% of reading or better; bidirectional • Pipe materials: steel, stainless steel, cast iron, PVC, copper • Fluid temperature: −10 to +150 °C • Data logging; analog (4–20 mA) or digital output; transducers IP65 • Calibration certificate shall cover converter and supplied transducers; mounting accessories and coupling gel shall be included
15 | Ultrasonic thickness gauge | • Range (steel): 1–300 mm (or wider); resolution 0.1 mm • Accuracy: ± 0.5% + 0.05 mm • Programmable sound velocity for different materials • Dual-element probe ≥ 5 MHz included with coupling gel and reference block • Delivered calibrated for thickness measurement; reference block included for routine verification
* - a measuring tool that is specially required for conducting energy audits in industrial enterprises
** - minimum required measurement tools for conducting energy audits in residential buildings