## Tabelle1
Specifications
No. | Item | Specifications |  | Use
1 | Total Organic Carbon Analyzer with Sample Feeder | Principle | Analysis of organic carbon in water by converting the organic carbon in the sample to carbon dioxide (CO₂), followed by quantitative analysis of CO₂ to determine the TOC content.. | Laboratory equipment for ONEA used to accurately measure the concentration of total organic carbon (TOC) in water, supporting water quality monitoring.
|  | Option | Automatic sample dilution, automatic acid addition, and automatic gas purge.
|  | Oxidation Technology | high-temperature catalytic oxidation (680°C - 1000 °C)
|  | Détector | NDIR, high sensitivity and stability.
|  | Feeder | Automatic sampler
|  | Analysis Parameters: | TC, TIC, TOC, NPOC
|  | Operating Gas Requirement | The system must be compatible with Oxygen, nitrogen or purified air as a carrier/combustion gas Note : Gas cylinders are not to be supplied with the equipment.
|  | Sample Type | Water sample
|  | Measurement Range | 0,05mg/l to 1000 mg/l (automatic dilution up to 100.000 mg/l)
|  | User interface | PC (PC with installed software and connecting cable must be included in delivery)
|  | Software | Functionality: Control of the analyzer and sample feeder; Storage of calibration curves, data management, custom reports. Software must support english or french language.
|  | Detection Limit | ≤ 0.05 mg/L Carbon
|  | Maximum Allowed Error | ≤ 5 %
|  | Injection Volume | Adjustable, within 10 - 1000µl range
|  | Power Supply | Voltage 220 V ± 10%, 50/60 Hz (standard); Overvoltage protection included
|  | Accessories | Must be included: Required cables, hoses, adapters, calibration solutions, maintenance kit including spare fuses.
|  | Documentation | user manual in french or english
|  | Warranty | at least 2 years and according to GIZ General Purchase Conditions 2023
2 | Flame Photometer | Parameters | Analysis of Sodium (Na+), Potassium (K+), Calcium (Ca2+), Lithium (Li+), or Barium (Ba++) | Laboratory equipment for ONEA. Used to measure the concentration of certain chemical elements in water.
|  | Fuel required: | Propane or butane gas
|  | Air supply source : | Internal or external compressor (moisture and oil free)
|  | Power supply : | 220 V ± 10% at 50 or 60Hz automatic selection; Overvoltage protection
|  | Measurement method | Filters for the analysis of different defined elements
|  | Detection limit (LOD) and quantification limit (LOQ) | Na: ≤ 0.1 ppm; K:≤ 0.05 ppm; Li: ≤0.05 ppm; Ca: ≤ 0.5 ppm; Ba: ≤1.6 ppm
|  | Sample feeder | Automatic feeder with at least 40 positions, automatic sample dilution; Automatic introduction of blanks and standards; Automatic detection of the last sample, stop in case of flame extinction.
|  | Safety | Gas leak alarms, automatic gas cut-off, protection against accidental extinction
|  | Connectivity | USB port for data download
|  | User Interface : | TFT or LCD screen
|  | Advanced functions | Storage of calibration curves, automatic calculation of the correlation coefficient.
|  | Acessories | Required accessories must be included: cables, gas hose, gas adapter, desktop computer, calibration solutions, spare fuses, an additional nebulizer, maintenance kit),
|  | Documentation | user manual in french or english
|  | Warranty | at least 2 years and according to GIZ General Purchase Conditions 2023
3 | Automatic Kjeldahl Nitrogen Analyser | A) Automatic Kjeldahl Analyzer |  | Laboratory equipment for ONEA. This equipment enables accurate determination of total Kjeldahl nitrogen (TKN) in water samples
|  | Principle | System integrating automatic cleaning, high precision feed pumps, and titration. Real-time control of distilled liquid temperature
|  | Measurement range: | 0.1 mg - 200 mg N
|  | Reproducibility: | ≤ 0.5%
|  | Recovery: | 99.5% - 100%
|  | Power supply: | 220 VAC ±10%, 50/60 Hz
|  | Electrical Protection : | Integrated protection against overvoltage, power surges, and short circuits.
B) Kjeldahl Digestion System
|  | Batch capacity | 20 to 40 samples per batch
|  | Digestion tube capacity | 250ml to 300 ml
|  | Material | Digestion chamber must be corrosion-resistant against acid
System for pre-treating samples orremoving acids after digestion
|  | Heating of digestion chamber: | Adjustable from ambiant up to 430 °C; precision ± 1°C to allow for complete digestion of the sample. Automatic temperature control. Thermal insulation to ensure thermal stability and user protection
|  | Power supply: | 220 V ±10%, 50/60 Hz
|  | Electrical Protection : | Integrated protection against overvoltage, power surges, and short circuits.
C) Suction System
|  | Principle | Adjustable vacuum pressure (-0.02 to -00.8MPa); stable transfer system
|  | Material | Pipes and pump must be corrosion resistant to acid
|  | Filtration | Ternary filtration system (water condensation, alkali neutralization, and active carbon filtration)
|  | Documentation | user manual in french or english
|  | Warranty | at least 2 years and according to GIZ General Purchase Conditions 2023