## Neonatal Platform
Tender document
Specification | [5 empty] | Entry by the bidder
Item No | Specification | Unit | Quantity | Installation and Commissioning (Yes/No) | Training (Yes / No) | (Please confirm the specification with "Yes" and tell us the brand name, model, country of origin and other information about the offered product)
1 | Open Resuscitation System for Newborns (Neonatal nursing platform) | Piece | 1 | Yes | Yes
1.1 | General requirements for the product: Open neonatal resuscitation platform suitable for neonatal resuscitation, stabilization, monitoring, and thermoregulation of newborns including premature infants in delivery rooms, operating theatres, and NICU settings. The system shall comply with applicable IEC/ISO standards for neonatal medical equipment and shall be suitable for safe operation in hospital environments. Availability of spare parts minimum 5 years after procurement.
1.2 | Dimensions: approx 1200 mm x 650 mm x1800 mm (LxWxH)
1.3 | Weight: approx 120 kg
1.4 | Main display: color LCD touch screen, approxately 10 inch
1.5 | Heating system: Radiant infrared heater
1.6 | Heating modes: Automatic and manual modes must be available
1.7 | Phototherapy module
1.8 | Temperature monitoring: Continuous skin temperature monitoring with servo control
1.9 | Temperature alarms: Audible and visual alarm system for temperature deviation and sensor failure
1.10 | Resuscitation system: Integrated T-piece resuscitator for CPAP and PPV
1.11 | Pressure adjustment: Adjustable PIP and PEEP controls
1.12 | Pressure monitoring: Integrated airway pressure manometer or digital pressure display
1.13 | Oxygen blender: Adjustable FiO₂ range between 21–100% must be covered
1.14 | Gas supply compatibility: Compatible with centralized oxygen and compressed air supply systems
1.15 | Integrated neonatal monitoring
1.16 | Battery backup. Rechargeable lithium-ion battery or equivalent technology. Minimum 1 hour of continuous operation for monitoring and essential functions during mains power failure. Visual battery status and charging indication on the display or control panel.
1.17 | X-ray tray
1.18 | APGAR timer
1.19 | Advanced alarm system
1.20 | Integrated suction system
1.21 | Data storage/connectivity
1.22 | Weighing scale
1.23 | Respiratory circuit compatibility: Compatible with standard disposable neonatal breathing circuits
1.24 | Bed surface: Radiolucent mattress platform
1.25 | Bed positioning: Trendelenburg and reverse Trendelenburg adjustment
1.26 | Working height adjustment: Ergonomically adjustable working height
1.27 | Safety rails: Integrated side safety barriers
1.28 | Lighting system: Integrated LED examination light
1.29 | Lighting adjustment: Electronically adjustable illumination intensity
1.30 | User interface: Digital control panel with visual display
1.31 | Display parameters: Temperature, heater status, alarms, and operating modes
1.32 | System diagnostics: Automatic self-test during startup
1.33 | Material requirements: Smooth rounded edges and medical-grade biocompatible materials
1.34 | Mobility: Mounted on lockable antistatic castors
1.35 | Electrical data: Supply voltage 220–240 V
1.36 | Mains frequency: 50/60 Hz
1.37 | Safety system: Overheating protection and alarm system for power failure and sensor malfunction
1.38 | Cleaning and disinfection: Surfaces resistant to commonly used hospital disinfectants
2 | Accessories and Integrated Components:
2.1 | Integrated neonatal T-piece resuscitation system
2.2 | Skin temperature probes/sensors | Piece | 5
2.3 | Integrated airway pressure manometer
2.4 | LED examination light
2.5 | Air and oxygen connection hoses | Piece | 2
2.6 | Accessory tray or holders
2.7 | Neonatal mattress suitable for intensive care use
2.8 | Reusable or disposable neonatal breathing circuits | Piece | 5
2.9 | Power cable with Euro plug
3 | Information regarding installation The Open Neonatal Resuscitation and Stabilization platform will be installed at the Simulation Center of Center for Development of Professional Qualification of Medical Workers (CDPQMW) in Tashkent, Uzbekistan. The equipment must be handed over to the customer in a ready-to-use condition". Acceptance Protocol will be signed by Supplier and CDPQMW. The supplier shall provide all necessary accessories, cables, hoses, probes, and connectors required for complete installation and operation of the neonatal resuscitation system.
4 | Details regarding training: After intallation of the Open Neonatal Resuscitation and Stabilization platform at the Simulation Center of CDPQMW, authorized specialists by the manufacturer or the contractor have to conduct training. CDPQMW address: 51, Parkent street, Tashkent, Uzbekistan Training Objectives: Upon completion of this initial user training course, participants will be able to: - Use the neonatal resuscitation system properly and safely in delivery room and NICU settings. - Operate all standard and advanced functions of the neonatal resuscitation platform, including radiant warming, thermoregulation, monitoring, and T-piece resuscitation functions. - Correctly use the integrated T-piece resuscitator, oxygen/air blender, CPAP and positive pressure ventilation functions. - Understand and configure temperature control modes, alarm systems, APGAR timer, and optional monitoring functions. - Safely operate optional configurations such as phototherapy, weighing scale, suction system, and patient monitoring modules. - Perform routine cleaning, disinfection, functional checks, maintenance, and safe handling procedures according to manufacturer recommendations. - Recognize alarm conditions and apply appropriate troubleshooting and safety measures. - Perform basic preventive maintenance and daily care procedures correctly and safely. Number of participants: Up to 4 neonatal healthcare professionals, biomedical engineers, technical specialists, or clinical instructors. Requirements on Participants: Participants should have basic experience in neonatal care, neonatal resuscitation, intensive care equipment, or biomedical equipment operation and maintenance. Requirements on Trainers: Authorized specialists by the manufacturer or the contractor Duration of training: Minimum 3 training days of 8 hours each. Language of the training: The training shall be conducted in Russian or Uzbek. If the trainer can only speak English, GIZ will provide translators for Russian or Uzbek.
5 | Details regarding transportation and delivery of equipment: Terms of Delivery: DPU Tashkent. Before shipping of goods, the proforma-invoice should be provided to GIZ country office to check whether all data correspond and are correct for the subsequent customs clearance.
General Information
Project location, climate and other information: Tashkent is the capital city of Uzbekistan. The city is located in northeastern Uzbekistan and serves as the political, economic, scientific, and transportation center of the country. According to official data, Tashkent has a population of approximately 3 million inhabitants. The city is situated at an altitude of about 440–480 meters above sea level. Tashkent is connected to national and international transport networks through Tashkent International Airport, railway infrastructure, and major road corridors. The climate in Tashkent is continental with hot, dry summers and relatively cold winters. The average annual temperature is approximately 16°C, with seasonal temperatures ranging from around -5°C in winter to +40°C in summer. The average annual rainfall is approximately 440 mm, mainly occurring during winter and spring months. Average annual humidity ranges between 50% and 60%. These climatic conditions should be considered for transportation, storage, installation, and operation of medical equipment. In some areas, temporary interruptions of the electrical power supply may occasionally occur. Therefore, medical equipment with integrated battery backup systems, voltage protection, or compatibility with emergency power supply systems is recommended to ensure continuous and safe operation.
Mandatory requirements for all material items
Standards Unless otherwise specified herein, all materials used, equipment supplied and service to be provided shall conform to the latest applicable editions of ISO, IEC, EN, DIN, or equivalent standards, which include but not limited to: ISO 13485 (Medical devices – Quality management systems) (the certificate must be included in the submitted offer) IEC 60601 series for medical electrical equipment safety IEC 60601-1-8 for alarm systems IEC 60601-2-21 or equivalent standards applicable to infant radiant warmers and neonatal care equipment Applicable MDR/MDD or equivalent regulatory requirements for medical devices (the certificate or evidence must be included in the submitted offer)
Material and equipment requirements All material must be new, current in distribution (not a discontinued model if the successor model is already in distribution). All material must be fully inspected and tested to demonstrate compliance with specification requirements. (Manufacturer's Declaration of Conformity and Product Certificate). Electrical equipment requirements: Compatible with 220–240 V AC power supply with grounding conductor Compatible with 50/60 Hz frequency systems Equipment shall tolerate recurrent unexpected interruptions of the power supply without damage to electrical, electronic, monitoring, or mechanical components Integrated battery backup or compatibility with emergency power systems is recommended due to possible temporary power interruptions Additional requirements: All measurements, displays, indicators, and controls shall conform to the metric system Materials and components shall be suitable for operation in continental climates with hot summers and varying humidity conditions Surfaces shall be resistant to commonly used hospital cleaning agents and disinfectants All patient-contact materials shall be biocompatible and suitable for neonatal use Labels, safety markings, and symbols shall be provided in English and/or local language with internationally understandable symbols
The documentation shall include, but not be limited to (shall be supplied in English and/or Russian language): General description of the equipment and accessories Spare parts lists and/or spare parts catalogues and ordering information User manuals and technical documentation, minimum two copies Detailed documentation for transportation, storage, installation (Installation drawings and connection requirements), operation, cleaning, disinfection, troubleshooting, preventive maintenance, and servicing of the equipment Technical specifications and safety precautions Information regarding battery operation and emergency procedures Recommended service intervals and calibration procedures