## LOT1
LOT | Pos. | Item | Technical Specification | Qty | Fully compliant
Y / N
1 | 1.1 | RADIONE Ebola Detection Kit (RP017) | Automated POCT-PCR | 15
Sample Type: Human serum or plasma
Target patient: Human
Target: Ebola virus (Bundibugyo, Zaire, Sudan, Tai Forest, Reston)
| 1.2 | RADIONE 4ch (KM010) | Throughput: 4 samples | 3
Kit type: Extraction + PCR
Technology: Magnetic rod extraction + PCR
Sample volume: <500µL
Processing time: ≤65 minutes
Automatic UV Sterilization: Yes

## LOT2
LOT | Pos. | Item | Technical Specification | Qty | Fully compliant
Y / N
2 | 2.1 | Glove Box | General Requirements • Glove Box providing maximum personnel, product, and environmental protection. • Fully enclosed, gas-tight construction. • The glove box meets the technical and safety standards of glove boxes adequate to BSL-3 and BSl-4 laboratory enviroments. • Compliance with WHO Laboratory Biosafety Manual recommendations and applicable EN 12469 and/or NSF/ANSI 49 standards. • Portable or Semiportable | 3
Airflow and Pressure Regulation • Continuous negative pressure: The chamber must maintain internal negative pressure to ensure air only flows inward if a leak occurs. • One-way airflow: Air supply must be dedicated and controlled, moving exclusively from the outside environment into the cabinet through filters. • Dedicated exhaust system independent of room ventilation.
Filtration and Containment • Supply air filtration: via HEPA filter (minimum H14 or equivalent efficiency), or comparable. • Double HEPA exhaust: Internal air must pass through two HEPA filters (or comparable) installed in a series before being discharged outside. • Filter efficiency ≥99.995% at MPPS or equivalent certification standard. • No air recirculation: The system must operate on a 100% exhaust cycle; no contaminated air is allowed to recirculate inside.
Barriers and Material Transfer • Hermetic sealing: The main body must be gas-tight and typically constructed from heavy-duty, chemical-resistant stainless steel. • Double-door pass box: Materials and samples must enter and exit through antechambers equipped with interlocking doors to prevent simultaneous opening. • Decontamination integration: Transfer ports must feature integrated sterilization mechanisms, such as chemical dunk tanks, autoclaves, or pass-through UV/vapor systems.
Gloves and Ergonomics • Heavy-duty, chemical-resistant long gloves: Arm-length gloves must be mechanically attached to the glove ports with airtight seals. • Safe-change glove ports: The port design must allow operators to replace worn gloves from the outside without breaching containment or losing negative pressure. • Two (not four) glove ports
Monitoring and Safety Alarms • Audible and/or visual alarms: The unit must trigger immediate warnings if negative pressure drops or if airflow falls below safe operational limits. • Backup power compatibility: Control systems and exhaust fans must becapable of being connected to the emergency power supply to maintain containment during blackouts.