## Tabelle1
Specifications
No. | LOTs | Item | Specifications |  | Use | Quantity
[6 empty] | amount | unit
1 | Lot1 | Total Organic Carbon Analyzer with Sample Feeder | Principle | Analysis of organic carbon in water by converting the organic carbon in the sample to carbon dioxide (CO₂), followed by quantitative analysis of CO₂ to determine the TOC content.. | Laboratory equipment for ONEA used to accurately measure the concentration of total organic carbon (TOC) in water, supporting water quality monitoring. | 1 | pieces
|  |  | Option | Automatic sample dilution, automatic acid addition, and automatic gas purge.
|  |  | Oxidation Technology | high-temperature catalytic oxidation
|  |  | Détector | NDIR, high sensitivity and stability.
|  |  | Feeder | Automatic sampler
|  |  | Analysis Parameters: | TC, TIC, TOC, NPOC
|  |  | Operating Gas Requirement | The system must be compatible with Oxygen, nitrogen or purified air as a carrier/combustion gas Note : Gas cylinders are not to be supplied with the equipment.
|  |  | Sample Type | Water sample
|  |  | Measurement Range | 0-1000 mg/l automatic dilution of concentrated samples to extend measuting range to 100.000 mg/l
|  |  | User interface | PC (PC with installed software and connecting cable must be included in delivery)
|  |  | Software | Functionality: Control of the analyzer and sample feeder; Storage of calibration curves, data management, custom reports. Software must support english or french language.
|  |  | Detection Limit | 0.05 mg/L Carbon
|  |  | Maximum Allowed Error | ± 5 %
|  |  | Injection Volume | Variable
|  |  | Power Supply | Voltage 220 V ± 10%, 50/60 Hz (standard); Protection against electrical overvoltage
|  |  | Accessories | Must be included: Required cables, hoses, adapters, calibration solutions, maintenance kit including spare fuses.
|  |  | Documentation | user manual in french or english, Remote support in french or english must be available to help/advise during the installation of the system if needed
2 | Lot2 | Flame Photometer | Parameters | Analysis of Sodium (Na+), Potassium (K+), Calcium (Ca2+), Lithium (Li+), or Barium (Ba++) | Laboratory equipment for ONEA. Used to measure the concentration of certain chemical elements in water. | 1 | pieces
|  |  | Fuel required: | Propane or butane gas
|  |  | Air supply source : | Internal or external compressor (moisture and oil free)
|  |  | Power supply : | 220V AC at 50 or 60Hz automatic selection; Overvoltage protection
|  |  | Measurement method | Filters for the analysis of different defined elements
|  |  | Detection limit (LOD) | Na ≤ 0.1 ppm; K ≤ 0.05 ppm; Li ≤ 0.02 ppm; Ca: ≤0.2 ppm; Ba: ≤20 ppm
|  |  | Sample feeder | Automatic feeder with at least 40 positions, automatic sample dilution; Automatic introduction of blanks and standards; Automatic detection of the last sample, stop in case of flame extinction.
|  |  | Safety | Gas leak alarms, automatic gas cut-off, protection against accidental extinction
|  |  | Connectivity | RS-232 and/or USB for data download (a USB–RS-232 adapter is acceptable)
|  |  | User Interface : | LCD screen or LED
|  |  | Advanced functions | Storage of calibration curves, automatic calculation of the correlation coefficient.
|  |  | Acessories | Required accessories must be included: cables, gas hose, gas adapter, desktop computer, calibration solutions, spare fuses, an additional nebulizer, maintenance kit),
|  |  | Documentation | user manual in french or english, Remote support in french or english must be available to help/advise during the installation of the system if needed
3 | Lot 3 | Automatic Kjeldahl Nitrogen Analyser | A) Automatic Kjeldahl Analyzer |  | Laboratory equipment for ONEA. This equipment enables accurate determination of total Kjeldahl nitrogen (TKN) in water samples | 1 | pieces
|  |  | Principle | System integrating automatic cleaning, high precision feed pumps, and titration. Real-time control of distilled liquid temperature
|  |  | Measurement range: | 0.1 mg - 200 mg N
|  |  | Reproducibility: | ±0.5%
|  |  | Recovery: | ≥99.5%
|  |  | Power supply: | 220 VAC ±10%, 50 Hz
|  |  | Electrical Protection : | Integrated protection against overvoltage, power surges, and short circuits.
|  |  | B) Kjeldahl Digestion System |  |  | 1 | Pieces
|  |  | Batch capacity | At least 20 samples in one batch
|  |  | Digestion tube capacity | At least 300 ml
|  |  | Material | Digestion chamber must be corrosion-resistant against acid
System for pre-treating samples orremoving acids after digestion
|  |  | Heating of digestion chamber: | Up to 430 °C to allow for complete digestion of the sample. Automatic temperature control. Thermal insulation to ensure thermal stability and user protection
|  |  | Power supply: | 220 VAC ±10%, 50 Hz
|  |  | Electrical Protection : | Integrated protection against overvoltage, power surges, and short circuits.
|  |  | C) Suction System |  |  | 1 | Pieces
|  |  | Principle | Fluid suction provided by a vacuum pump with adjustable pressure, guaranteeing controlled and stable transfer to the analysis system
|  |  | Material | Pipes and pump must be corrosion resistant to acid
|  |  | Filtration | Ternary filtration system (water condensation, alkali neutralization, and active carbon filtration)
|  |  | Documentation | user manual in french or english, Remote support in french or english must be available to help/advise during the installation of the system if needed