## Leistungsbeschreibung
Kriterienkatalog
Allgemeines
A control solution for up to 32 individual feedback loops. The control boxes must include the ability for both fast and slow feedback for driving piezo electric transducers and Peltier elements, respectively. The control system must be remotely controllable and must include the ability to tailor the proportional-integral-differential (PID) controller.
Hauptlieferort:
Universität Paderborn
Warburger Straße 100
33098 Paderborn
Kriterienkatalog
All requirements for the delivery and service are defined in the following criteria catalog.
Explanations of the criterion types in column C:
A: Exclusion criterion (non-fulfillment leads to exclusion from the offer).
B: Evaluation criterion; in column F you can see how many points you get depending on the degree of fulfillment for an evaluation criterion.
I: Information required (not included in the evaluation).
column A | column B | column C | column D | column E | column F | column G
Response YES / NO
Pos. | Product and material requirements, certifications and standards - valid for all offered products. | Criteria type | YES | NO | Achievable score and remarks | Information of the bidder
0 | The price for the complete system including all services such as delivery and installation does not exceed 135.700,-€ (excl. VAT). | A | o | o
1 | Complete System Requirements
1.1 | A complete turnkey system including a remote user interface. | A | o | o
1.2 | The electrical connections must be single-phase 220-240 V, 50 Hz. | A | o | o
1.3 | The equipment specifications must be capable of operating under laboratory conditions at a room temperature of 21°C +/- 2°C and a humidity of 45% +/-20%. | A | o | o
1.4 | The complete system consists of a sufficient number of control electronics (see Section 2) for controlling at least 32 feedback loops individually | A | o | o
1.5 | The control electronics can address at least 32 Peltier elements and at least 32 Piezo elements | A | o | o
1.6 | The complete system can be mounted in 19" racks without further modification | A | o | o
1.7 | The complete system must be CE certified | A | o | o
1.8 | The system can be controlled via one or more touchscreens | A | o | o
1.9 | The system has at least RS-232 and USB connectivity | A | o | o
1.10 | The system comes with all required accessories such as cables (at least 5m long) and adapters | A | o | o
|  | A | o | o
2 | Core Specifications of Electronics
2.1 | The system provides at least 32 fast servo controls that each are as specified in 2.1.x | A | o | o
2.1.1 | The control has an operating temperature from 5 to 40 degrees | A | o | o
2.1.2 | The control has an analogue input voltage range from -6 V to +6 V | A | o | o
2.1.3 | The control has an RF level monitor range from -30 dBm to +10 dBm | A | o | o
2.1.4 | The control has the following RF level monitor accuracy in the frequency range from 10 MHz to 100 MHz: min -1 dBm, max 1dBm with an Output level differential at no load of +/- 27 V | A | o | o
2.1.5 | The control has a main output, monitor main output, and auxiliary output | A | o | o
2.1.6 | The control has a monitor for the aux output and a monitor for an error signal | A | o | o
2.2 | The system provides at least 32 temperature controls that each are as specified in 2.2.x | A | o | o
2.2.1 | The control has an operating temperature from 15 to 45 degrees | A | o | o
2.2.2 | The control has a temperature sensor range from 12.5 degree C (or less) to 39.5 degree C (or more) | A | o | o
2.2.3 | The control has a temperature sensor absolute accuracy of +/- 0.5 degree C | A | o | o
2.2.4 | The control has a 15-way D-SUB-HD connector for the Heater/Peltier | A | o | o
2.2.5 | The control has an operating voltage from -5 V (or less) to +5 V (or more) | A | o | o
3 | Control and Software
3.1 | If the manufacturer provides a PC, the software must run stable on the delivered PC | A | o | o
3.2 | If the manufacturer does not provide a PC, the software must run stable on windows 11 | A | o | o
3.3 | All required functions to operate the detector system must be operable from the provided software | A | o | o
3.4 | The system must be controlable through programming/scripting languages such as e.g. Python, Labview, Matlab | A | o | o
4 | Delivery and Payment
4.1 | Delivery DPU Paderborn (Incoterms 2020) The offer must include delivery to Universität Paderborn, PhoQS, P10, 33098 Paderborn and installation. | A | o | o
4.2 | Delivery and installation must be completed before March 2026 | A | o | o | Indicate the delivery time in column G
6 | Other
6.1 | System performance result datasheet and certifications are included | A | o | o
6.2 | The manufacturer offers support in both German and English | A | o | o
6.3 | The bidder must explain how he will ensure technical support in the long term (> 10 years). | A | o | o
6.4 | On standard working days, the manufacturer must be reachable by phone at any time between 9:00 a.m. and 3:30 p.m. The service is provided by a service branch in Europe, with English-speaking service employees. | A | o | o
6.5 | The supplier must confirm that it will respond to inquiries by telephone or email within 3 (three) working days. | A | o | o
Date, Signature and Stamp